Lab Name:      					Procedure #:

Procedure: CLIA Complexity: WAIVED (Urine) CLIA Complexity: MODERATE (Serum)

| Prepared By | Date Adopted | Supersedes Procedure # |
| --- | --- | --- |

| Review Date | Revision Date | Signature |
| --- | --- | --- |

| Distributed to | # of Copies | Distributed to | # of Copies |
| --- | --- | --- | --- |

This Procedural Bulletin is intended to provide a ready outline reference for performance of the assay. It is the obligation of every manufacturer of medical devices labeled FOR IN VITRO DIAGNOSTIC USE to provide a complete package insert in accordance with FDA labeling regulation (21 CFR 809.10). Prepared in accordance with the guidelines recommended by the Clinical and Laboratory Standards Institute, Wayne, PA 19087; CLSI Document GP2-A2.

IMCO provides CLSI procedures for your use.  The procedures are required to include the same information as listed in the package insert. Any modifications to this document are the sole responsibility of the Laboratory

CLIA Complexity: WAIVED

#### INTENDED USE

The IMCO One Step HCG Urine/Serum Test in vitro diagnostic test is used for the qualitative determination of hCG in urine or serum for early detection of pregnancy. It detects the presence of human chorionic gonadotropin (“HCG”) at a concentration of 5mlU/ml or above. The IMCO One Step HCG Urine/Serum Test is intended for professional use only (Clinical Laboratory Use).

#### PRINCIPLE

Human Chorionic Gonadotropin (HCG) is one of the glycoproteins secreted by the placenta. It contains α-and β-subunits. The α-subunit of HCG shows cross-reactions with LH, TSH and FSH, because of similarities in the amino acid sequences. However, the β-subunit of HCG is specific, and shows the lowest cross-reactions in assays. The antibodies, which were used before in HCG detection, were directed against the whole HCG molecule, and therefore regularly presented cross-reactions with LH and TSH and FSH. One Step HCG test is aimed to be a HCG specific test by using a monoclonal antibody against highly purified β-HCG.

The kit is a rapid qualitative assay and based on the principle of double antibodies sandwich immunoassay for determination of HCG in urine or serum. When specimen is added into the sample well,

Lab Name:      					Procedure #:

the urine/serum migrates upward by chromatographic, HCG will be captured by β-HCG Monoclonal antibody in sample pad and then meet α-HCG Monoclonal antibody on the T line and Anti-mouse IgG on the C line coated in the membrane, if the HCG is at or above 25 mlU/ml there is red bands will be appeared both in T line and C line which indicates positive result; If the HCG is zero or below 25 mlU/ml there will no red band appear in the T line, only a red band in C line which indicates negative result. To serve as a procedure control, a colored line will appear at the C line, if the test has been performed properly.

#### REQUIRED MATERIAL - PROVIDED

- 25 tests per kit, one test in one pouch.

- One pouch containing a reaction cassette, a dropper and a desiccant. The desiccant is for storage purposes only, and is not used in the test procedures.

- One Package Insert.

#### REQUIRED MATERIAL – NOT PROVIDED

- A clean, dry, plastic or glass container to collect the urine.

- A clean tube to collect the serum

- Timer (watch or clock)

- Control materials

#### PRECATIONS

- This kit is for external use only. Do not swallow.

- Discard after first use. The test strip cannot be used more than once.

- Do not use test kit beyond expiration date.

- Do not use the kit if the pouch is punctured or not well sealed.

- Keep out of the reach of children.

#### STORAGE AND STABILITY

- Store at 4ºC to 30ºC in the sealed pouch up to the expiration date.

- Keep away from direct sunlight, moisture and heat.

- DO NOT FREEZE.

- Preferably open the pouch only shortly before the test.

#### SPECIMEN COLLECTION AND PREPARATION

Urine:

Any urine specimen is appropriate for Pregnancy Testing but the first morning urine specimen is optimal because of its highest concentration of HCG. Urine specimens exhibiting visible precipitates should be centrifuged, filtered, or allowed to settle to obtain a clear specimen for testing.

Serum:
Collect blood aseptically by venipuncture into a clean tube without coagulants, separate the serum from blood as soon as possible to avoid hemolysis.  Use clear non-hemolyzed specimens only.

#### SPECIMEN STORAGE

Specimens may be stored at 2-8°C for up to 48 hours prior to testing. For longer storage, specimens may be frozen and stored below -20°C. Frozen specimens should be thawed, thoroughly mixed, and brought to room temperature before testing.

Lab Name:      					Procedure #:

#### DIRECTIONS FOR USE

- The test cassette and urine/serum specimen has to be at room temperature (18-30°C) for testing.

- Remove a Testing Device from the foil pouch by tearing at the notch and place it on a level surface.

- Holding a sample dropper vertically, add exactly four drops of the urine/serum specimen to the sample well (with an arrow marked).

- Read results between 3-5 minutes. Do not read results after 5 minutes.

#### INTERPRETATION OF RESULTS

Negative (not pregnant):

Only one color band appears, in the Control Zone. No apparent band on the Test Zone. This indicates that no pregnancy has been detected.

Positive (pregnant):

Distinct color bands appear in the Control and Test Zones. It indicates that you are pregnant. The color intensity of the test bands may vary since different stages of pregnancy have different concentrations of HCG hormone.

Invalid:

No visible band at all or there is a visible band only in the test region and not in the control region. Repeat with a new test kit. If test still fails, please contact IMCO Technical Support at 1-844-369-2344 for further assistance.

NOTE: if the color in the Test Zone is weak, it is recommended that the test be repeated in 48 hours.

#### QUALITY CONTROL

- A red line appearing in the control region (C) confirms that sufficient specimen volume has been added.

- User must follow the federal, state, and local law requirements for quality control testing.

- A positive and a negative hCG controls should be used to verify proper test performance.

#### LIMITATIONS

- As it is with any diagnostic procedure, a confirmed pregnancy diagnosis should only be made by a physician after evaluating all clinical and laboratory findings.

- Elevated levels of hCG can be caused by a few conditions other than pregnancy. Therefore, the presence of hCG in urine/serum specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.

Lab Name:      					Procedure #:

- A normal pregnancy cannot be distinguished from an ectopic pregnancy based on hCG levels alone. Also, spontaneous miscarriage may cause confusion in interpreting test results.

- Sterility treatments, based on hCG, may cause false results. Consult your physician.

- If a negative result is obtained with a urine specimen and pregnancy is still suspected, obtain a first morning urine specimen and re-test. The serum must be clear non-hemolysis.

- A negative result from a urine or serum specimen collected from a woman in very early pregnancy may be due to an unusually low concentration of hCG. In such cases, the test should be repeated on a fresh specimen obtained approximately two days later.

- When hCG is greater than 100 lU/mL it may present negative result because of hook effect.

- Heterophile antibodies, human anti-animal antibodies (such as human anti-mouse antibody-HAMA), rheumatoid factor and autoantibodies may lead to falsely elevated hCG results.

#### EXPECTED RESULTS

Negative results are expected in non-pregnant women. Although the hCG levels in normal early pregnant women are variable, IMCO One Step HCG Urine/Serum Test is capable of detecting pregnancy as early as 1 day after the first missed menses.

#### PERFORMANCE CHARACTERISTICS

#### ACCURACY

The result from IMCO One Step HCG Urine/Serum Pregnancy Test is compared with that from a commercially available hCG test. Among 582 urine specimens, IMCO One Step HCG Urine/Serum Pregnancy Test shows 283 positive and 299 negative while the comparison reagent shows 282 positive and 300 negative. Among 582 serum specimens, IMCO One Step HCG Urine/Serum Pregnancy Test shows 282 positive and 300 negative while the comparison reagent shows 282 positive and 300 negative.  In the study, IMCO One Step HCG Urine/Serum Pregnancy Test showed > 99% overall agreement when compared with other urine membrane HCG test.

#### OTHER COMMERCIAL METHODS

Urine specimen:

|  | IMCO One Step HCG Urine/Serum Pregnancy Test |
| --- | --- |
|  | ＋ | － | Total |
| Comparison reagent | + | 281 | 1 | 282 |
|  | － | 2 | 298 | 300 |
|  | Total | 283 | 299 | 582 |

Serum specimen:

|  | IMCO One Step HCG Urine/Serum Pregnancy Test |
| --- | --- |
|  | ＋ | － | Total |
| Comparison reagent | + | 280 | 2 | 282 |
|  | － | 2 | 298 | 300 |
|  | Total | 282 | 300 | 582 |

Agreement of the positive with comparison reagent：>99 %

Agreement of the Negative with comparison reagent :> 99%

Total Agreement with comparison reagent：> 99%

Lab Name:      					Procedure #:

#### SENSITIVITY AND SPECIFICITY

The test limit of IMCO One Step HCG Urine/Serum Test is 25 mIU/mL or greater. The test has been standardized to the W.H.O. Standard. No cross-reaction appears when LH (300 mIU/mL), FSH (300 mIU/mL), and TSH (1,000 mIU/mL) were added to the negative (0 mIU/mL hCG) and positive (25 mIU/mL hCG) specimens.

Urine samples:

Serum samples:

| Results Concentration (mIU/mL) | Positive results | Negative results | Total number of samples |
| --- | --- | --- | --- |
|  | + | - |  |
| 0 | 0 | 5 | 5 |
| 12.5 | 0 | 5 | 5 |
| 18.75 | 0 | 5 | 5 |
| 25 | 29 | 1 | 30 |
| 50 | 10 | 0 | 10 |
| 100 | 5 | 0 | 5 |
| Total | 60 | 60 |

#### INTERFERING SUBSTANCE

The following potentially interfering substances were added to hCG negative and positive specimens.

Urine specimen:

| Acetaminophen | 20mg/dl | Hemoglobin | 20mg/dl |
| --- | --- | --- | --- |
| Acetylsalicylic acid | 20mg/dl | Tetracycline | 20mg/dl |
| Ascorbic acid | 20mg/dl | Ampiciline | 20mg/dl |
| Atropine | 20mg/dl | Albumin | 20mg/dl |
| Caffeine | 20mg/dl | Bilirubin | 2mg/dl |
| Centesic acid | 20mg/dl | Glucose | 2mg/dl |

Lab Name:      					Procedure #:

Serum specimen:

| Acetaminophen | 20mg/dl | Hemoglobin | 20mg/dl |
| --- | --- | --- | --- |
| Acetylsalicylic acid | 20mg/dl | Tetracycline | 20mg/dl |
| Ascorbic acid | 20mg/dl | Ampicilline | 20mg/dl |
| Atropine | 20mg/dl | Albumin | 20mg/dl |
| Caffeine | 20mg/dl | Bilirubin | 40 mg/dL |
| Centesic acid | 20mg/dl | Thiglycerides | 1200mg/dL |
| Glucose | 2mg/dl |  |  |

None of the substances at the concentration tested interfered in the assay.

#### LINEARITY REPORTABLE RANGE

Considerating a hook effection in immunoassay, we investigate HCG assay range. Spike high HCG concentration solution to negative urines from female or male, in spiked samples, the concentration of HCG is from 500IU/ml to 10 mIU/ml, each sample be tested repeatedly.

The result showed the test can detect the HCG concentration at range 25mIU/ml—100IU/ml. If HCG level above 100IU/ml, a hook effect may be observed.

#### BIBLIOGRAPHY

- Batzer FR. “Hormonal evaluation of early pregnancy”, Fertil. Steril. 1980; 34(1): 1-13

- Braunstein GD, J Rasor, H. Danzer, D Adler, ME Wade “Serum human chorionic gonadotropin levels throughout normal pregnancy”, Am. J. Obstet. Gynecol. 1976; 126(6): 678-681

- Braunstein GD, JL Vaitukaitis, PP Carbone, GT Ross “Ectopic production of human chorionic gonadotropin by neoplasms”, Ann. Intern Med. 1973; 78(1): 39-45

- Catt KJ, ML Dufau, JL Vaitukaitis “Appearance of hCG in pregnancy plasma following the initiation of implantation of the blastocyte”, J. Clin. Endocrinol. Metab. 1975; 40(3): 537-540

- Lenton EA, LM Neal, R Sulaiman “Plasma concentration of human chorionic gonadotropin from the time of implantation until the second week of pregnancy”, Fertil. Steril. 1982; 37(6): 773-778

- Rasor, J.L. Braunstein G.D., 1977., Obstet. Gynecol. 50, 553-558.

Manufactured for:

Independent Medical Co-op Inc.

Daytona Beach, FL-32114

www.imcoinc.com

Technical Support:1-844-369-2344

Lab Name:      					Procedure #:

| LOG SHEET |
| --- |
| Record Built-in Procedural Controls on the first patient tested each day. |
|  | Date | Patient Name | Positive Procedural Control | Test Results Read at 3 minutes Do not read the test results after 5 (Five) minutes | Lot Number and Exp. Date | Technician |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |

Lab Name:      					Procedure #:

| LOG SHEET |
| --- |
|  | Date | External Positive Control Result | External Negative Control Result | Lot Number and Exp. Date | Technician |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |
| 5 |  |  |  |  |  |
| 6 |  |  |  |  |  |
| 7 |  |  |  |  |  |
| 8 |  |  |  |  |  |
| 9 |  |  |  |  |  |
| 10 |  |  |  |  |  |
| 11 |  |  |  |  |  |
| 12 |  |  |  |  |  |
| 13 |  |  |  |  |  |
| 14 |  |  |  |  |  |
| 15 |  |  |  |  |  |

# Text boxes

| Results Concentration (mIU/mL) | Positive results | Negative results | Total number of samples |
| --- | --- | --- | --- |
|  | + | - |  |
| 0 | 0 | 5 | 5 |
| 12.5 | 0 | 5 | 5 |
| 18.75 | 0 | 5 | 5 |
| 25 | 29 | 1 | 30 |
| 50 | 10 | 0 | 10 |
| 100 | 5 | 0 | 5 |
| Total | 60 | 60 |